Evaluation of the Degradation of Materials by Exposure to germicide UV-C Light through Colorimetry, Tensile Strength and Surface Microstructure Analyses

Supplementary Material

S1. Conversion relation of degradation time and real experiment time

To carry out the aging analysis of materials, the Zenzoe lamp from ASTI MOBILE ROBOTICS and Boos Technical Lighting was located into a closed room with no windows. This lamp integrates 5 fluorescent tubes of UV radiation in the C band, emitting at a wavelength of 254 nm and with a power of 140 W. The lamp was placed in a test configuration where samples could be irradiated placing them in different shelves with different distances to the lamp (Figure S1).

Figure S1. Two different perspectives of the distances of the room where the aging tests of the materials were carried out. The lamp was located in the middle of the room, having shelves were the samples could be placed around it.

In order to determine if the radiation reached all the point homogeneously, radiation measurements were performed with a UV-C sensor placed in the same position and with the same orientation that the samples would have (Figure S2). The description of all the positions of the samples can be found in Table S1.

Table S1. Materials irradiated with UV-C light.

| Position | Location (looking from the lamp) | Measured mean value (μW/cm2) | Measured maximum value (μW/cm2) |
| --- | --- | --- | --- |
| 1 | 103 cm from the left corner of the shelf | 1218 | 1550.5 |
| 2 | 8 cm from the left corner of the shelf | 759.6 | 809.3 |
| 3 | 37 cm from the right corner of the shelf | 1134.8 | 1141.9 |
| 4 | 37 cm from the right corner of the shelf | 931.9 | 962.7 |
| 5 | 103 cm from the left corner of the shelf | 1033.3 | 1623.8 |
| 6 | 8 cm from the left corner of the shelf | 675.9 | 773.8 |
| 7 | 8 cm from the right corner of the shelf | 1231.9 | 1272.5 |
| 8 | 8 cm from the right corner of the shelf | 1522.5 | 1576.2 |
| 9 | 30 cm from the left corner of the shelf | 1273.6 | 1325.7 |
| 10 | 30 cm from the left corner of the shelf | 1168.2 | 1186.6 |

Figure S2. Location of the UV-C sensors for the radiation measurement.

Once knowing the dose that the robot generated and having measured the radiation at each one of the points where the samples could be placed, an equivalence between the use of the robot once a day for every day of the year and the radiation reached during the tests needed to be estimated. This provided the opportunity to analyze the aging effect of long periods of a “once a day” exposure (i.e. years), with shorter prolonged exposure times (i.e. hours).

If the samples were placed vertically and parallel to the lamp, the dose levels would be 36.8 mJ/cm2 for each use at a speed of 100 mm/s. Therefore, the equivalent dose values for each position depending on the years treated would be the ones shown in Table S2. The approximate time of equivalence is about 3 hours for each year of use. Also, it is worth mentioning that the radiation process is cumulative, so two treatments a day for a year would be the same as 2 years with a daily pass.

Table S2. Equivalence between the use of the robot for a whole year once a day every day and the hours that a sample has to be exposed continuously in each of the positions in the closed room at Ceit.

| Years treated once a day, every day, at 100 mm/s | Accumulated dose (mJ/cm2) | Equivalent hours for each of the positions |
| --- | --- | --- |
|  |  | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 |
| 1 year | 13432 | 3.1 | 4.9 | 3.3 | 4.0 | 3.6 | 5.5 | 3.0 | 2.5 | 2.9 | 3.2 |
| 2 years | 26864 | 6.1 | 9.8 | 6.6 | 8.0 | 7.2 | 11.0 | 6.1 | 4.9 | 5.9 | 6.4 |
| 3 years | 40296 | 9.2 | 14.7 | 9.9 | 12.0 | 10.8 | 16.6 | 9.1 | 7.4 | 8.8 | 9.6 |
| 4 years | 53728 | 12.3 | 19.6 | 13.2 | 16.0 | 14.4 | 22.1 | 12.1 | 9.8 | 11.7 | 12.8 |
| 5 years | 67160 | 15.3 | 24.6 | 16.4 | 20.0 | 18.1 | 27.6 | 15.1 | 12.3 | 14.6 | 16.0 |
| 6 years | 80592 | 18.4 | 29.5 | 19.7 | 24.0 | 21.7 | 33.1 | 18.2 | 14.7 | 17.6 | 19.2 |
| 7 years | 94024 | 21.4 | 34.4 | 23.0 | 28.0 | 25.3 | 38.6 | 21.2 | 17.2 | 20.5 | 22.4 |
| 8 years | 107456 | 24.5 | 39.3 | 26.3 | 32.0 | 28.9 | 44.2 | 24.2 | 19.6 | 23.4 | 25.6 |
| 9 years | 120888 | 27.6 | 44.2 | 29.6 | 36.0 | 32.5 | 49.7 | 27.3 | 22.1 | 26.4 | 28.7 |
| 10 years | 134320 | 30.6 | 49.1 | 32.9 | 40.0 | 36.1 | 55.2 | 30.3 | 24.5 | 29.3 | 31.9 |
| 11 years | 147752 | 33.7 | 54.0 | 36.2 | 44.0 | 39.7 | 60.7 | 33.3 | 27.0 | 32.2 | 35.1 |
| 12 years | 161184 | 36.8 | 58.9 | 39.5 | 48.0 | 43.3 | 66.2 | 36.3 | 29.4 | 35.2 | 38.3 |
| 13 years | 174616 | 39.8 | 63.9 | 42.7 | 52.0 | 46.9 | 71.8 | 39.4 | 31.9 | 38.1 | 41.5 |
| 14 years | 188048 | 42.9 | 68.8 | 46.0 | 56.1 | 50.6 | 77.3 | 42.4 | 34.3 | 41.0 | 44.7 |
| 15 years | 201480 | 45.9 | 73.7 | 49.3 | 60.1 | 54.2 | 82.8 | 45.4 | 36.8 | 43.9 | 47.9 |
| 16 years | 214912 | 49.0 | 78.6 | 52.6 | 64.1 | 57.8 | 88.3 | 48.5 | 39.2 | 46.9 | 51.1 |
| 17 years | 228344 | 52.1 | 83.5 | 55.9 | 68.1 | 61.4 | 93.8 | 51.5 | 41.7 | 49.8 | 54.3 |
| 18 years | 241776 | 55.1 | 88.4 | 59.2 | 72.1 | 65.0 | 99.4 | 54.5 | 44.1 | 52.7 | 57.5 |
| 19 years | 255208 | 58.2 | 93.3 | 62.5 | 76.1 | 68.6 | 104.9 | 57.5 | 46.6 | 55.7 | 60.7 |
| 20 years | 268640 | 61.3 | 98.2 | 65.8 | 80.1 | 72.2 | 110.4 | 60.6 | 49.0 | 58.6 | 63.9 |

S2. Tensile strength test samples

In this section, images of all the different material samples used in this study are shown before and after the tensile strength test. This way, it can be observed how the samples ended after performing this assay.

Figure S3. Different materials before the performance of the tensile strength test. The samples shown in this image are the control samples (the ones that were not exposed to any UV-C light).

Figure S4. Different materials after the performance of the tensile strength test for all the controls and samples exposed for the equivalent dose of UV-C light exposure times. The times indicated in the image (1, 2, 4, 8 and 16 years) are the equivalent time to daily short irradiations, which in continuous treatment during the experiments were 3.5, 7, 14, 28, and 54 hours of exposure.